# Gravity: Analog ORP Sensor Meter for Arduino

In the ever-increasing need and awareness for health and safety, water has become a focus point for proper disinfection treatment as aftereffects of industrialization can become a danger for contamination of water via direct contact. Water is a critical control point of contact where proper treatment can greatly reduce these microbes, thereby reducing the chances of illness and pollution.

        ORP SENSOR

An Oxidation Reduction Potential (ORP) Sensor measures the activity of oxidizers and reducers in an aqueous solution. It is a potentiometric measurement from a two-electrode system similar to a pH sensor. Sometimes it is also referred to as a redox measurement. Unlike a pH sensor, an ORP sensor measures the ratio of oxidized to reduced forms of all chemical species in solution.
The ORP sensor is made up of two electrochemical half cells where the reference electrode is generally Ag/AgCl and the measurement electrode is commonly Pt. The potential difference between the two electrodes represents the redox potential of the solution being measured and can be described by the Nernst equation.

E = Eo – 2.3 (RT/nF) x (log [Ox] / [Red])

 Where:
E = total potential developed between the measurement and reference electrodes
Eo = a voltage specific to the system
R = gas constant
 T = temperature in K
n = the number of electrons involved in the equilibrium between the oxidized and reduced species
F = Faraday constant
[Ox] = concentration of the oxidized species
[Red] = concentration of the reduced form of that species

        Hydroponic ORP

Hydroponic growers are used to using pH and EC meters to control their growing conditions but very few use ORP meters in order to learn more about their nutrient solution. An ORP meter can give you very useful information and cheap ORP meters can usually be bought on ebay or amazon for less than 20 dollars each. Today we will talk about ORP meters, what they are, what they are useful for and how you can use them in your hydroponic crop.
–

–
An ORP meter or “Oxidation-Reduction Potential” meter characterizes the electrochemical environment within a solution. An ORP meter has two electrodes, a traditional reference electrode with a known potential and a platinum electrode whose potential changes depending on how hard the solution is pushing either to take electrons from the platinum electron or give it electrons. A solution that has a lot of substances that are willing to give electrons will tend to give a more negative reading and a solution that has more substances willing to take electrons will tend to give a more positive reading. The ORP reading is given in mV.
So how is this useful in hydroponics? It is useful in the sense that we can know exactly how the chemical environment is behaving. The ORP of potable water is generally around 600-700mV, this is because oxidants – substances that are willing to take electrons – are added to solutions in order to kill pathogens. The chemical environment needs to have an ORP of above 600mV to eliminate harmful fungal spores and bacteria. Of course this means that if you want to run a sterile hydroponic environment you’ll want to keep the ORP of your solution probably in the 300-500mV range, large enough to prevent any micro-organisms from growing but low enough to prevent any damage from happening to your roots.
In this way you can use things like hypochlorous acid and hydrogen peroxide to increase the “killing power” of your solution while knowing how harsh you’re making the chemical environment. The ORP will also give you signs about water oxygenation and biological activity within the water. A reductive environment – ORP below 100mV – will mean that there is a significant number of substances in the solution that want to give electrons and these substances are generally organic acids, bacteria, viruses or other organics molecules, like reductive sugars. If this is the case then it means that oxygen in solution has a short lifetime so you will want to increase your oxygenation significantly or your roots might be starved of this essential nutrient.
–

–
The chemical environment is determined by the sorts of half reactions that can happen in solution and this is also determined by the pH of your solution. The above table shows some of the most common electrochemical half-reactions that can happen in solution. For example in order to reduce molecular oxygen and obtain 4 electrons we need to produce hydroxide ions. This means that oxidation reactions will tend to increase the pH and therefore they are expected to become harder as the pH rises. We also have the opposite case for hydrogen peroxide where a more acidic solution is bound to prevent the oxidation of peroxide to molecular oxygen. It is worth noting that these are half reactions so in reality what always happens is that two half-reactions – for example oxygen reduction and Fe oxidation – are brought together to generate a chemical change in the environment.
In the end the ORP measurement gives you something that pH and EC measurements do not tell you, which is what the chemical environment looks like from an oxidation-reduction perspective. With this information it becomes easier to tell things like whether you’re lacking enough oxygenation, whether you’re adding too much hypochlorite or peroxide and whether or not you should be adding more or less microbes to your environment.

# Gravity: Analog ORP Sensor Meter for Arduino
(https://wiki.dfrobot.com/Gravity_Analog_ORP_Sensor_PRO_SKU_SEN0464 )
In the ever-increasing need and awareness for health and safety, water has become a focus point for proper disinfection treatment as aftereffects of industrialization can become a danger for contamination of water via direct contact. Water is a critical control point of contact where proper treatment can greatly reduce these microbes, thereby reducing the chances of illness and pollution.
While traditional ORP meters are hard to get or use, this Analog ORP Meter is a highly compact ORP monitoring system that fits into your Arduino. This design configuration allows the user to accurately monitor ORP without having to add any additional circuitry or components to your design. Communication with the Analog ORP Meter is done using only simple wiring and programming. You may also check Liquid Sensor Selection Guide to get better familiar with our liquid sensor series.
ORP (Oxidation-Reduction Potential) is a measure of the ability of oxidation and reduction of the aqueous solution, characterization of oxidizing or reducing the relative degree. Unlike a pH measurement that follows a logarithmic curve and therefore requires more calibration adjustments, ORP follows a linear relationship and does not need instrument adjustment as much as it needs electrode maintenance. ORP has proven to be a reliable method of measuring water quality and provides the operator with a single value of measurement regardless of which product, commodity, operation, treatment, or sanitizer is used, and regardless of varying field conditions or methods by which chemical is applied.
Attention：
The bottle cap of the probe contains protective liquid (3.3mol/L KCL). Even if the bottle cap is screwed tightly, some protective liquid may still leak around the bottle cap, forming white crystals. But as long as there is still protective liquid in the bottle cap, it will not affect the life and accuracy of the probe. So please use it with confidence. It is recommended that the white crystals be poured back into the protective liquid in the bottle cap.
#### Applications
- Water Quality Testing
- Aquaculture
- Marine Science
- Hydroponic Gardening
- General Chemistry

#### Specification
- Module Power:+5.00V
- Module Size: 40mmX27mm(1.57"x1.06")
- Measuring Range:-2000mV—2000mV
- Suitable Temperature:5-70℃
- Accuracy:±10mv (25 ℃)
- Response Time:≤20sec
- ORP Probe with BNC Connector
- PH2.0 Interface(3 foot patch)
- Zero calibration button
- Power Indicator LED

ORP Composite Electrode Manual

ORP (Oxidation Reduction Potential) composite electrode is made up of gold or platinum
electrode and reference electrode. It is used to measure the oxidation reduction potential of
the solution.
2. Main Characteristics
In the ORP standard solution, the millivolt value is about 222±15mV(25℃).
3. Test the Electrode in the ORP Standard Solution
Let the ORP electrode connected to a pH/mV meter, and set the electronic unit on the mV
measurement.
Immerse the ORP composite electrode in the standard solution, stirring it period of time,
and then place it statically. You can only read the value after the reading is stable.

Use and Maintenance

When use the ORP electrode, the calibration is not necessary. It can be directly
connected to the instrument. Only when you doubt the quality or the test results, you
can use a standard solution to test the mV value, as a method of distinguishing
electrode quality.

2) There are reasonable electrode soaking solutions in the bottle, the insert electrode tip is
soaking in it. Before measurement, unscrew the cap and pull out the electrode, use the
pure water to clean the electrode and dry it before use. Please do not rub the surface of
the sensitive elements. Before measurement, the electrode should be put in the
solution and stirred, in order to acceleration response. After measurement, the
electrode should be washed in the pure water, pulled into the bottle and screw the cap,
in order to prevent the leakage. You should exchange the solution when it is Turbid or
moldy.

3) The configuration of the electrode soaking solution:
Take a packet of buffer whose pH value is 4.00, dissolved it in 250mL water. Then add
56g muriate of potash (analytically pure) in together, then heating and stirring them
until completely dissolved.

4) Cleaning and activation of the ORP electrode:
After a long-term use, the pollution on the sensitive elements will lead to the inaccurate
measurement and the slow response. You can clean and activate the electrode as
follows:
1)For inorganic pollution, put the electrode into 0.1mol / L dilute hydrochloric acid
for about 30 minutes, then clean it with pure water. After the cleaning, put it in the
electrode soaking solution for about 6 hours

2) For organic oil or oil film pollution, you can wash the sensitive elements with the
   detergent, and then wash it cleanlily with the pure water. After the cleaning, put it
in the electrode soaking solution for about 6 hours.

        3)When the pollution on the sensitive elements is serious, forming an oxide or
reduction membrane on the surface, the finer metallographic sandpaper is
appropriately used to polish the surface of the sensitive elements. After the polish,
put the electrode in the electrode soaking solution for about 6 hours.